## Tier system_v02-00
Supplemental Table I | [5 empty] | -- required metadata --
Category | Tier Nr. | Name | Description | Example Experiment Type | Example Labelling | Experimental/ Sample | Microscope hardware specifications | Image acquisition settings | Optical calibration | Intensity calibration | Mechanical calibration
Descriptive | 1 | Minimum Information/ Material & Methods | Reporting qualitative effects, or effects that require simple quantification including the identification of non-refractive limited objects followed by basic feature extraction and statistical analysis | Transfection control, viability assay, counting of cells and nuclei, expression level measurements, localization of markers in cellular sub-compartments | Fluorescent In Situ Hybridization (FISH), Immuno Fluorescence (IF), Fluoresent Protein (FP) labelling | experimenter name; experiment description and date; sample description; mounting medium; temperature and CO2 conditions; | microscope manufacturer, model and type; light source manufacturer, wavelenght and type; objective manufacturer, magnification, NA and correction; filter/dichroic transmittance range; detector manufacturer and type | acquisition date; immersion liquid name and refractive index; illumination type and intensity; fluorophore; exposure time; pixel dwell time; channel name, color, contrast method and acquisition mode; image dimension order and number; physical pixel size x, y, and z | not required; recommended quarterly | not required; recommended annually
Analytical | 2 | Advanced Quantification and/or Live Cell Imaging | Identification and localization of diffraction-limited particles, super-resolution microscopy, tracking of intracellular dynamics | Diffraction-limited spot localization, measurement of distances, co-localization studies, signal-starved features, advanced processing, cell tracking and single-particle tracking, dynamic expression level quantification | All of the above + Single Molecule (SM) FISH, CasFISH, SM Proximity Ligation Assay (PLA), dCas9-based labelling, OligoPaint | O2 pressure, and humidity conditions; refractive index of the mounting medium; thickness of the coverglass | detailed environmental control device, microscope table, light source, light source coupling, transmittance light path, magnification, sample positioning, focusing, autofocus, filter, dichroic, additional optics and detector specification (e.g., lightsource spectral properties; objective correction properties; focusing device ZReproducibility, ZSettlingTime, ZResolution, etc.) | illumination attenuation; objective temperature and iris aperture; immersion liquid measured refractive index; sample positioning settings; detector integration; ligthpath configuration | required monthly | highly-recommended monthly to quarterly
| 3 | Manufacturing/ Technical Development/ Full Documentation | Full documentation of microscopic setup, image acquisition and quality control | Microscopy hardware manufacturing; development of novel unproven technology in both commercial and academic settings; full reproducibility of microscopy set-up and image acquisition settings | All of the above | all the metadata specified by the data model - including any novel technology-specific metrics |  |  | required for every acquisition | required monthly to quarterly

## Tier system_v02-00 SUMMARY
Supplemental Table I
Category | Tier Nr. | Name | Description | Example Experiment Type | Example Labelling | Experimental/ Sample | Microscope hardware specifications | Image acquisition settings | Quality Control
Descriptive | 1 | Minimum Information/ Material & Methods | Reporting qualitative effects, or effects that require simple quantification including the identification of non-refractive limited objects followed by basic feature extraction and statistical analysis | Transfection control, viability assay, counting of cells and nuclei, expression level measurements, localization of markers in cellular sub-compartments | Fluorescent In Situ Hybridization (FISH), Immuno Fluorescence (IF), Fluoresent Protein (FP) labelling | experimenter name; experiment description and date; sample description; mounting medium; temperature and CO2 conditions; | microscope manufacturer, model and type; light source manufacturer, wavelenght and type; objective manufacturer, magnification, NA and correction; filter/dichroic transmittance range; detector manufacturer and type | acquisition date; immersion liquid name and refractive index; illumination type and intensity; fluorophore; exposure time; pixel dwell time; channel name, color, contrast method and acquisition mode; image dimension order and number; physical pixel size x, y, and z | Not required or recommended quarterly to annualy depending on the type of procedure
Analytical | 2 | Advanced Quantification and/or Live Cell Imaging | Identification and localization of diffraction-limited particles, super-resolution microscopy, tracking of intracellular dynamics | Diffraction-limited spot localization, measurement of distances, co-localization studies, signal-starved features, advanced processing, cell tracking and single-particle tracking, dynamic expression level quantification | All of the above + Single Molecule (SM) FISH, CasFISH, SM Proximity Ligation Assay (PLA), dCas9-based labelling, OligoPaint | O2 pressure, and humidity conditions; refractive index of the mounting medium; thickness of the coverglass | detailed environmental control device, microscope table, light source, light source coupling, transmittance light path, magnification, sample positioning, focusing, autofocus, filter, dichroic, additional optics and detector specification (e.g., lightsource spectral properties; objective correction properties; focusing device ZReproducibility, ZSettlingTime, ZResolution, etc.) | illumination attenuation; objective temperature and iris aperture; immersion liquid measured refractive index; sample positioning settings; detector integration; ligthpath configuration | Highly recommended or required monthly to quarterly depending on the type of procedure
| 3 | Manufacturing/ Technical Development/ Full Documentation | Full documentation of microscopic setup, image acquisition and quality control | Microscopy hardware manufacturing; development of novel unproven technology in both commercial and academic settings; full reproducibility of microscopy set-up and image acquisition settings | All of the above | All the metadata specified by the guidelinesl - including any novel and technology-specific metrics |  |  | Required for every acquisition, monthly or quarterly depending on the type of procedure

## Tier system_v02-00 MINIMAL
Category | Tier Nr. | Name | Description | Example Experiment Type | Example Labelling | Microscope Hardware Specifications and Acquisition Settings | Quality Control
Descriptive | 1 | Minimum Information/ Material & Methods | Reporting qualitative effects, or effects that require simple quantification including the identification of objects larger than the resolution limit followed by basic feature extraction and statistical analysis | Transfection control, viability assay, counting of cells and nuclei, expression level measurements, localization of markers in cellular sub-compartments | Fluorescent In Situ Hybridization (FISH), Immuno Fluorescence (IF), Fluoresent Protein (FP) labelling | Manufactuer, model and type of microscope, lightsource, objective, filter, dichroic and detector; acquisition date, immersion liquid refractive index, exposure time, image dimension order and number; physical pixel size x, y, and z | Not required or recommended quarterly to annualy depending on the type of procedure
Analytical | 2 | Advanced Quantification and/or Live Cell Imaging | Identification and localization of diffraction-limited particles, super-resolution microscopy, tracking of intracellular dynamics | Diffraction-limited spot localization, measurement of distances, co-localization studies, signal-starved features, advanced processing, cell tracking and single-particle tracking, dynamic expression level quantification | All of the above + Single Molecule (SM) FISH, CasFISH, SM Proximity Ligation Assay (PLA), dCas9-based labelling, OligoPaint | Detailed environmental control device, microscope table, light source, light source coupling, transmittance light path, magnification, sample positioning, focusing, autofocus, filter, dichroic, additional optics and detector hardware specification (e.g., lightsource spectral properties; objective correction properties; focusing device ZReproducibility, ZSettlingTime, ZResolution, etc.) and settings (e.g., illumination attenuation; objective temperature and iris aperture; immersion liquid measured refractive index; sample positioning settings; detector integration; ligthpath configuration) | Highly recommended or required monthly to quarterly depending on the type of procedure
| 3 | Manufacturing/ Technical Development/ Full Documentation | Full documentation of microscopic setup, image acquisition and quality control | Microscopy hardware manufacturing; development of novel unproven technology in both commercial and academic settings; full reproducibility of microscopy set-up and image acquisition settings | All of the above | All the metadata specified by the guidelinesl - including any novel and technology-specific metrics | Required for every acquisition, monthly or quarterly depending on the type of procedure